## Slide 1: DXS Optics Simulation
Nan Wang, Apr 30th 2021

## Slide 2: Outline
Crystal heat-load deformation
  Deformation profile
  Beam profile (xz, yz central slices)
  Beam profile around focal plane (xy)
  Strehl ratio
HRM tolerance scan

## Slide 3: Deformation
Crystal deformation is simulated as the height error for a flat mirror. For a 250um (FWHM) beam profile, the RMS height error is ~88.7 pm at 70W.

## Slide 4: Beam profile vs height error

## Slide 5: Around focal plane, ±25cm
flat mirror
8.87pm
26.62pm

## Slide 6: Strehl ratio
The theoretical Strehl ratio for the three cases are: 1.0, 0.952, 0.643 respectively. The simulated value is smaller due to astigmatism. Working on correction with 1D CRL.
0.952

## Slide 7: Crystal-lens distance

## Slide 8: Crystal miscut error, C1, C2

## Slide 9: Crystal miscut error, C3, C4

## Slide 10: To do
Finish the angular tolerance scan
Then switch to parabolic mirrors and for future simulations also do mirrors
SRW takes out quadratic terms for the parabolic mirrors
Divergent lens