## Slide 1: DXS Optics Simulation
Nan Wang, Feb 24th 2021

## Slide 2: Outline
Create beamline using Sirepo
Add time domain simulation
Optimize sampling parameters
Simulation results
Issues

## Slide 3: Sirepo
1. define source type and input source parameters
2. define beamline by dragging optical elements to the definition area
  Note: the central photon energy for each optical element is not automatically updated. The default photon energy is 9keV.
  Also, it would be nice if the radius of curvature for the CRLs can be calculated from focal distance.

## Slide 4: Python scripts from Sirepo
The python script generated by Sirepo is structured in the following way:
  Import modules
  Function for defining beamline optics, set_optics
  List of parameters, varParam
  Function for beam propagation, main()
The main benefit for using Sirepo is that we only need little modifications to the beamline optics and list of parameters when doing parameters scans.

## Slide 5: Adding time domain simulation
In the list of parameters:
  In Data Folder:
    Change ‘gbm_st’, ‘f’, rms pulse duration
  In Calculation Types:
    Change ‘w_e’, ‘f’, time_window_min
    Change ‘w_ef’, ‘f’, time_window_max
    Change ‘w_ne’, ‘i’, # of points in time/energy
    Change ‘w_smpf’, ‘f’, sampling factor (disabled if <0)
    Change ‘w_u’, ‘i’, 2, the correct electric field units
  A sample of modified code can be seen on the right, note that code for other source types can be commented out/deleted since they are not used.

## Slide 6: Adding time domain simulation
In the propagation function:

## Slide 7: Optimize parameters
Spatial: adjust the range scaling factor in the Propagation parameters at targeted optical elements so the entire beam profile is enclosed. Then adjust the spatial sampling resolution until aliasing disappears.
Temporal: first adjust ‘w_ne’ to make sure the incident pulse profile is sampled smoothly. Then use the resize function when propagating the beam when needed.
100fs, 9481eV, one asymmetric reflection (Si 444, 29.5 degree)

## Slide 8: Results
Typical simulation results with open and closed slits (10 um). The beamline is shown below

## Slide 9: Crystal-lens distance tolerance
Changing the distance between C2-CRL1 and CRL2-C3
100fs
200fs
400fs

## Slide 10: Issues
For a 4 crystal pulse stretcher/compressor using Si 111 with 5 degree asymmetry angles in the +-+- or -+-+ configuration, the output pulse durations don’t seem to be that different

## Slide 11: Issues, output
+-+-
-+-+