## Slide 1: DXS Optics Simulation Summary
Nan Wang, Jun 22nd 2021

## Slide 2: Timeline
SRW, lens
  HRM only
  HHLM only
  Telescope only
SRW, mirror
Lcls-beamline-toolbox
  Full beamline
  ShapeError correction

## Slide 3: SRW, HRM Only, focus
The vertical dimension is suppressed to 8 pts.
20fs incident Gaussian beam.
Can see some aliasing in log scale

## Slide 4: SRW, HRM Only, output
Output bandwidth is about 2.2 meV.

## Slide 5: SRW, HHLM Only, output

## Slide 6: SRW, HHLM+HRM, focus
Can see similar aliasing as in the HRM only case.
The closed slit case cannot be simulated due to sampling limitations.

## Slide 7: SRW, HRM Only, focus, mirror
20fs incident Gaussian beam.
Plots shown here are in linear scales. However, we can see similar aliasing as in the lens cases in log scale.

## Slide 8: SRW, HRM Only, output, mirror
We get very similar results compared to the lens case

## Slide 9: SRW, Telescope Only
SRW needs to update features including convex mirrors. So far, the shape of the mirrors are calculated theoretically and superimposed onto flat mirrors as shape error profiles in SRW.
Tried propagating SASE pulses through but eventually gave up due to tricky sampling along the beamline.

## Slide 10: LCLS, Full beamline, focus
open
closed

## Slide 11: LCLS, Full beamline, output
open
closed
closed

## Slide 12: ShapeError, HHLM1, 73.15W
SASE (option 1, 2), 2FWHM: 24.62 x 0.66mm
In any case, 70+W is too large. I don’t think we can correct for this at all.

## Slide 13: ShapeError, HHLM2, 50.1W
Self-seeded (option 3) , 2FWHM: 17.73 x 0.66mm
50W is still too large

## Slide 14: ShapeError, HHLM3, 1.78W
Self-seeded (option 3), 2FWHM: 17.73 x 0.66mm

## Slide 15: ShapeError, HHLM4, 1.69W
Self-seeded (option 3), 2FWHM: 24.62 x 0.66mm

## Slide 16: ShapeError, HRM1, 1.61W
Self-seeded (option 3), 2FWHM: 1.84 x 0.66mm

## Slide 17: ShapeError, HRM2, 1.88W
Self-seeded (option 3), 2FWHM: 1.48 x 0.66mm

## Slide 18: ShapeError, full, SASE
SASE (option 1)
Weakest link: HHLM1

## Slide 19: ShapeError, full, SASE (reduced at HRM)
SASE with reduced intensity for the HRM (option 2)
Weakest link: HHLM1

## Slide 20: ShapeError, full, self-seeded
Self-seeded (option 3)
Weakest link: HHLM2

## Slide 21: ShapeError, HHLM2, 50.1W, halved
Self-seeded (option 3) , 2FWHM: 17.73 x 0.66mm
Reducing the shapeError by a factor of 2 for HHLM2

## Slide 22: ShapeError, full, self-seeded, halved
Self-seeded (option 3)
Reducing the shapeError by a factor of 2 for all crystals
Plot to be updated

## Slide 23: ShapeError, HHLM2, 16.53W
SASE (option 1, 2), 2FWHM: 17.73 x 0.66mm
In the SASE case, the beam power at HHLM2 is reduced from 50.1W to 16.5W.
Plot to be updated